## Shipping Form
Please use one Shipping Form sheet per mail in holder. For multiple holders you should duplicate the Shipping Form sheet (right click on the Shipping Form worksheet tab, choose "Move or Copy", check the "make a copy" option and select the Shipping Form sheet from the list). See the instructions tab for more information and instructions. When you have finished, print out the Shipping Label sheet for labelling your parcel. Some of the cells have help boxes associated that you will get by clicking. Cells in red generally need to be filled in and check that the content is sensible.
B21 SAXS Mail In
| Visit ID: |  | enter your visit ID |  |  | Visit IDs should take the form i.e. mx12346-6
| Holder ID: |  | SAXS- | 1
| Name: |  | enter your name
| Contact e-mail: |  | enter your email
| Contact Phone (Optional): |  | enter a contact number
| Local Contact |  | CHOOSE A LOCAL CONTACT
| Storage Temperature |  | -80°C
HPLC Mode
| Holder ID | Sample position | Unique sample name | Flow rate (ml/min) | Buffer identifier and pH | Concentration (mg/ml) | Volume (ul) | Please select one column from the list | Notes
| B21-00001 | 1/A | my_sample_1 | 0.16 |  | 10 | 55 | Shodex KW-403
| B21-00001 | 2 |  | #N/A
| B21-00001 | 3 |  | #N/A
| B21-00001 | 4 |  | #N/A
| B21-00001 | 5 |  | #N/A
| B21-00001 | 6 |  | #N/A
| B21-00001 | 7 |  | #N/A
| B21-00001 | 8/H |  | #N/A
| B21-00001 | 9 |  | #N/A
Batch Mode
| Holder ID | Sample position | Sample Name | Buffer (True/False/In P9) | Buffer | Concentration | Volume (must be more than 25ul) | Molecular Weight (kDa) | Notes
| B21-00001 | 1/A | my_sample_1 | FALSE | 9
| B21-00001 | 2
| B21-00001 | 3
| B21-00001 | 4
| B21-00001 | 5
| B21-00001 | 6
| B21-00001 | 7
| B21-00001 | 8/H
| B21-00001 | 9 | my_buffer_1 | TRUE

## Instructions
Mail-in Instructions
Mail-in samples MUST be sent to the beamline in our bar-coded sample holders, to get holders for your experiment please contact your local contact by email (contact details below) letting them know how many samples you have to send and a mailing address to send the holders to. The sample holders have space for 8x 200 ul PCR tubes in a strip plus a 1.5 ml tube that is typically used to contain a matching buffer for batch mode experiments. Tubes can be most easily removed by pushing them up from underneath with a finger. Holes in the side of the holder are for standing the PCR strip up in to make pipetting easier. The shipping form contains all of the information that beamline staff need to measure your samples in the way that is appropriate, please fill out all of the relevant fields to avoid delays on the day of your beamtime. Please print off the shipping label from the Shipping Label workbook of this document and attach it to your package or copy the information from there to your own label or courier shipping docket.
When samples arrive on site and before they are taken to the beamline for data collection you can choose for your samples to be stored at:
-80°C | 4°C | Room Temp
BATCH mode sample requirements
Batch mode samples are loaded directly from the PCR tubes to the sample cell for measurement, samples take around 2 minutes each to run and this method allows SAXS measurements with well defined sample contents that is suitable for systematically changing parameters or ligand concentrations and studying the effects. In batch mode it is relatively hard to assess polydispersity, there is often small amounts of aggregate in the sample and there is often buffer mismatched and this makes batch mode less suitable for measuring data that will be used for modelling. Samples should be between 35 and 100 ul each. It is typical to do a serial dilution from 2 -> 0.1 mg/ml this is a rough guide and will vary from sample to sample. It is critically important that each batch mode sample come with at least 100 ul of a well matched buffer, typically a dialysate, the flow through from a spin concentrator or some eluent from the last sizing column of the purification process.
SEC mode sample requirements
SEC-SAXS mode allows a good assessment of polydispersity across an elution peak, removes aggregate from samples at the time of measurement and come with a built in accurate buffer blank. Data from SEC-SAXS is typically of a higher quality and more suited to model against. In SEC-SAXS mode it is harder to control parameters such as protein or ligand concentration and measurement times are longer, around 35 minutes per sample, making this mode less appropriate for screening parameters. SEC-SAXS samples are diluted as the elute through the sizing column and so sample concentrations should be higher. We typically use 50 μl samples at somewhere between 2 and 10 mg/ml. You should supply 5x or 10x running buffer stocks in 50 ml tubes. Unless a buffer component is very scarce or expensive please supply plenty of running buffer for HPLC pump washing and column equilibration as well as running samples. You will typically select from a range of columns that we keep at the beamline rather than supplying your own. The columns we have are:
Columns | Resolution range | Size | Flow rates (ml/min) |  | Pressure limit
|  | column volume (ml) | recommended (32mins) | max rate | Bar
Shodex KW-402.5 | 5-100 kDa | 4.6 | 0.16 | 0.25 | 85
Shodex KW-403 | 10-700 kDa | 4.6 | 0.16 | 0.25 | 85
Shodex KW-404 | 30 kDa - 4MDa | 4.6 | 0.16 | 0.25 | 85
Shodex KW-405 | 40 kDa - 20 Mda | 4.6 | 0.16 | 0.25 | 85
Superdex 200 Increase 3.2 | 10 kDa - 600 kDa | 2.4 | 0.075 | 0.075 | 30
Superose 6 Increase 3.2 | 5 kDa - 5 Mda | 2.4 | 0.075 | 0.075 | 30
*void volume is ~1/3rd of column volume
Accessing your data after the experiment
To access your data after the experiment please follow the instructions from the Diamond web site ( https://www.diamond.ac.uk/Users/Experiment-at-Diamond/IT-User-Guide/Not-at-DLS/Retrieve-data.html). Your experiment folder will contain: the raw images (.nxs and .h5 files) in the top level and a 'scanlist.csv' file linking the cryptic file names with the descriptive sample names. Subfolder 'hplc' contains traces from the HPLC, 'processed' contains .dat files for each image taken, 'processing' contains pipeline, calibration and mask nxs files used by Dawn for primary data reduction and a pypline subfolder that contains averaged and subtracted .dat files for batch mode experiments. You may find it useful to download only the scanlist file, the hplc, processed and processing folders and leave the large .nxs and .h5.
Help and Support
Should you need further assistance please contact your allocated local contact:
Nathan Cowieson | nathan.cowieson@diamond.ac.uk |  | +44 (0) 1235 56 7504
Katsuaki Inoue | katsuaki.inoue@diamond.ac.uk |  | +44 (0) 1235 77 8863
Nikul Khunti | nikul.khunti@diamond.ac.uk |  | +44 (0) 1235 56 7464
Charlotte Edwards-Gayle | charlotte.edwards-gayle@diamond.ac.uk |  | +44 (0) 1235 77 8033

## Shipping Label
TO: CHOOSE A LOCAL CONTACT
Diamond Light Source Ltd
Harwell Science and Innovation Campus
Didcot
Oxfordshire
OX11 0DE
Visit ID: ENTER YOUR VISIT ID
When at Diamond please store at: -80°C
Sender: enter your name, enter your email, enter a contact number