SXT Protocol V1.2.3

If you make changes, before you save: please increment last digit for wording changes, middle digit for smaller protocol changes, first digit for major changes.

Quick starter checklist for B24 staff on day of data collection

At B24 synoptic:

Open ‘Front End Status’ → ensure all absorbers are out → take exit slits to 60x60microns Y and X gaps → PGM should have Energy 500eV; Cff 1.6 and PGM Control to show Grating Translation 800l/mm; Mirror Translation Nickel

At the TXM (if onsite):

Make sure both VBG and VM2 are open at the TXM

Check that GM1 and GM2 are at high vacuum

Check ‘Cold Sample Chuck’ temperature is below 114K

De-frost the external nozzle of the LN2 dewar injector

Add liquid nitrogen to the cold trap and cover

At the TXM PC: For start data back up

Start Data Transfer script (shortcut located on Desktop)

Follow the instructions, starting by entering your user visit. Note: user visits have the form: ‘bi12345-6’

This will have created a folder following the format “D:\Data\<year>\<user visit>\”, if it didn’t already exist, and will transfer any files or folders from that to the appropriate raw data folder on b24-storage.

When collecting data, it must be saved to the directory stated above or it will not be processed.

Minimise (but do not close) the application during data collection

To stop the script, press `ctrl+c` with the Cygwin64 window selected.

Data collection using XRM Data Explorer

- Start XRM Data Explorer (icon located on the Desktop) → allow it to make changes

- Double click on the XRM Explorer Python module (top left key panel) and minimise the panel (do not close)

- Start Cyberlink YouCam (located on the Desktop)

- Start Sample Loader (located on the Desktop) → minimise extra black panel → Scan samples → click yes when asked if you are sure → wait until Procedure Status reads ‘Idle’

- Manually enter the sample names in the editable boxes

Note: all sample positions should register a ‘Sample+Cover’ state. If there is a position designated as empty, then that sample is already loaded on the sample stage.

Unloading a sample

If there is a sample loaded press unload → inspect the process through Cyberlink YouCam → wait until Procedure Status reads ‘Idle’ -- > Close Cyberlink YouCam

Loading a sample

Click ‘Load sample’ and chose the appropriate carousel position → inspect the process through Cyberlink YouCam → wait until Procedure Status reads ‘Idle’ → Close Cyberlink YouCam → minimise the loader panel

Stage 1: VLM imaging

- Open the motion control window (joystick button)

CRITICAL: On image control, do not confuse the step tab with the requested value tab or the current position tab. If you are not sure where these are, please ask!

1A. VLM imaging - Brightfield

- Move the X-ray objective to a safe place

  - Zoneplate → Zoneplate Z → set step 5000 → move one step positive (+)

- Turn the sample towards the 20x white light objective

  - Sample → Sample Theta = Set to -60 deg

- Select camera (top panel camera tab = magnifying lens button)

  - Choose VLM 20x

- Centre grid

  - Motion Control (top panel cogs tab) → Sample → Go to Sample X=-100 → Go to Sample Y=800

- Start continuous acquisition (Acquisition Settings = gears button)

  - Mode continuous, 0.5 sec exposure

  - Source settings → tick transmission → start

- Use Motion Control → Sample → Sample Z to get the grid into focus

  - Start with 50 um steps, then gradually reduce to 10 um steps as you approach focus

  - Continue until you can see grid letters clearly (or substrate holes in grids boxes)

  - abort acquisition

  - close image (do not save)

- Collect a mosaic of the whole grid (gears button)

  - Mode mosaic

  - camera settings 0.5 sec

  - source settings transmission ticked

  - Sample settings pixel size 0.37 µm

  - Mosaic horizontal step should read 323.75 Vertical step 323.75

  - Mosaic Rows 7 Columns 7

  - Mosaic axes sample X and sample Y

  - Start acquisition

  - Click Yes to the current stage angle

  - Save in Local Disk D:\<user_directory>\<sample_specification>\VLM_grid_mosaic

  - To adjust image contrast: On the image histogram panel (bottom left) select ‘Full range’ and chose histogram region to suit (left click and drag in the histogram view)

- (Optional) Transfer ROIs from cryoSIM using GridSNAP (instructions in the next section).

- Annotate with Areas-of-Interest / AOIs

  - Ensure the slider at the bottom of the mosaic is at frame 1

  - Frame AOIs with box tool from the side panel attached to the current image

  - Use line tool to number if needed

  - Save mosaic as VLM_grid_mosaic_annotated

- Source settings → tick transmission → start

- Turn the sample towards the Xray objective

  - Sample → Sample Theta = Set to 0 deg

- Return the X-ray objective to its imaging position

  - Zoneplate → Zoneplate Z → set step 5000 → move one step negative (-)

CRITICAL: Make sure that Zoneplate Z is >-10,000 before turning sample theta to –60deg

CRITICAL: Make sure that on completion of VLM mapping the sample is returned to 0deg and Zoneplate Z is reading the working position value which will be close to -14,000 (+/-200) depended on the setup.

1B. Transfer ROIs from SIM brightfield with GridSNAP (Optional if cryoSIm data were previously acquired at B24)

Please note that in GridSNAP you can hover your mouse over elements in the user interface, including the display, to see tips on how to use each of them. There is also a User Guide that can be accessed from the Help menu in the help section that includes some troubleshooting.

- Open GridSNAP (icon located on the Desktop)

- Import files to GridSNAP

  - Open the import dialog File → Import Files

  - Select your images to import:

    - The target image is the VLM mosaic from X-ray. Import the pipeline converted TIFF file from the ‘processed’ directory of your visit (\\b24-storage\data\<year>\biXXXX-X).

    - The transform image is the SIM brightfield mosaic. Import the pipeline converted TIFF file from the ‘raw’ directory of your visit (\\b24-1storage\data\<year>\biXXXX-X). The file will normally have the suffix ‘.ome.tiff’.

    - Click Submit.

  - Once imported the SIM image should be superimposed on the VLM image. The lines of the grid squares should be tilted and spaced equivalently, but not fully aligned yet.

- Align the images

  - Look at the two grid images and select the rotation option (-90, 0, +90, 180) that moves the finder markings to the same side of the grid square.

  - If the finder markings are flipped (mirror image) between images (e.g. C4 is backwards on the VLM but not the SIM), select the Flip option.

  - Click and drag the SIM image until it is overlayed on the same area of the VLM image (you can hover the mouse over the display to see the controls). It may help to adjust the compositing options for this.

  - Fine tune your alignment, if needed. Due to focusing differences and the grid bending slightly between microscopes, this is never going to be perfect, so you should aim for the best compromise between areas. Repeat the following two steps until you are happy with it:

    - Adjust the scale and rotation sliders to improve the alignment.

    - Zoom in to an area and drag the SIM image to better align that section.

Note: you can change the colours and transparency of the SIM and X-ray brightfield mosaics if it aids in better alignment or visualisation.

- Use the transformed coordinates

  - Use the ROI coordinates from the table in GridSNAP to move to each AOI:  Motion Controller → Sample → Sample X/Y → Go to move to the motor this position. This must be repeated* for each ROI as you go through the remaining steps of data acquisition.

  - When you put the mouse over the display, the coordinates of the target (VLM) image will be displayed, which may be helpful if an ROI was not marked.

Note: depending on your grid and ROIs, you may want to zoom in to an area and adjust the SIM image slightly to optimise alignment for that area before copying the coordinates of those AOIs.

1C. VLM imaging – Fluorescence (Optional: depending on project requirements)

At the TXM (Note: These steps can only be done in person at the beamline)

- Plug in the LED light source (near the TXM vacuum chamber)

- Set the LED channel and filter wheel to the appropriate colour

  - For green fluorescence use LED channel 3, Filter 4

  - For red fluorescence use LED channel 2, Filter 1

- Make sure all view ports are covered - i.e. turn off the bright field lamp and cover the aperture with foil.

At the PC (back to Data Explorer)

- Start continuous acquisition (Acquisition Settings = gears button)

  - Mode continuous

  - Camera Settings 20s (5 to 60s, depending on sample)

  - Source settings → untick transmission

  - Start

  - Abort when happy with exposure (binning will increase your signal if weak)

- Collect a focal series to find focus

  - acquisition mode focal series

  - camera settings

    - Exposure 5s

    - binning 2

  - source settings

    - none chosen

  - focal series

    - total projections 10

    - images per projection 1

    - focal axis sample Z

    - range 10

  - Start and save as FS_<number>.txrm

Note: to see the series in real time tick the box next to the slider at the bottom of the image

  - when done move the slider to the best focus slice

  - position pointer in centre of view → right click → chose the go to position option and confirm

  - acquisition mode single, 10s exposure (or best estimate), binning 1

  - save image VLM_FL_<colour>_<GridLocation = LetterNumberLetter>

Note: The arrow at the top of the side panel will expand the image menu if desired

- Repeat for other channels If needed (no further focusing required)

At the TXM (To be done in person at the beamline)

- Unplug the LED light source (left-most cable at the back of the controller box)

Stage 2: X-ray imaging

- Recentre the theta axis

  - Sample → Sample Theta → set to 0 (zero) degrees → press Go

- Microscope tab (top left of the main panel) → Camera temperature → Pixis → change set temperature –65C → apply

CRITICAL: Confirm with beamline staff this is safe to do

- Select camera (top panel camera tab = magnifying lens button)

  - Choose Pixis 1x

2A. Reference Acquisition

- Select location

  - open VLM grid mosaic

  - Select area on the VLM grid that has no membrane or cells (make sure image slider is at 1)

  - Position pointer in centre of area → right click → chose the go to position option → change theta to 0 degrees → and Go

  - Acquisition mode continuous, 0.5s exposure, sample settings pixel tick, pixel size 0.01

- Acquire reference

  - acquisition mode average, 0.5s exposure, sample settings pixel 0.01, total projections 10, Images per. projection 1, collect reference none, Average on the fly should not be selected

  - Use name Ref_<date as yyyymmdd> _<time as hhmm>

  - Start acquisition (note the image is not updating unless the bottom panel box on the image series is ticked)

  - Process image (panel top left of main view), average series of Images, choose input, tick despeckle, start

  - Save averaged file without renaming

2B. Xray mosaic of ROIs

- Find ROI

  - Open focused VLM of ROI (brightfield or fluorescence)

  - Position pointer close to features of interest → right click → choose the go to position option → change theta to 0 degrees → and confirm

Note: To choose a starting area ensure you are not on the thickest area of a cell and not near a grid bar; stay close to the biological material

- Focus the objective (zoneplate)

  - Start acquisition (gears button)

    - Mode continuous, 0.5s exposure

  - Move Motion Controller → Sample → Sample Z to get sample into approximate focus (1 to 5μm steps)

  - Focus image on support film or cell features of interest

  - Acquire mosaic

  - Abort when in focus

  - Start acquisition (gears button)

    - Mode mosaic, Camera setting 0.5s, binning 1, Sample settings pixel size 0.01 µm,

    - Mode mosaic, mosaic → images per projections 1 → horizontal step

    - vertical step 9.46 → Row 7, columns 7, horizontal axis sample X, vertical axis sample Y

    - Save as xray_mosaic_<GridLocation = LetterNumberLetter>

  - Reposition and repeat if a shift in view is required

  - Reference mosaic

  - Go to image and select the arrow button (top left) to expand the metadata

  - Choose the reference tab → single reference browse, chose the correct file, apply single reference

2C. Tilt series acquisition

CRITICAL: Be extremely cautious when changing Sample Theta during X-ray data collection. Never exceed -55 to +55 degrees range (-55 =< theta <= +55).

- Centre on ROI

  - open X-ray mosaic of the area with marked selected FOV

  - Place pointer on desired centre → right click → chose the go to position option → change theta to 0 degrees → and confirm

- Find optimal exposure for projections

  - Acquisition settings: mode continuous, camera setting 0.5s, start

  - Check pixel counts on your image (10-30k on average is good); change exposure in increments of 0.5sec to increase count if needed

  - Acquisition settings: Abort and delete image

- Move to tilt axis

  - Acquisition settings: mode continuous, camera setting 0.5s, start

  - Motion Control → Sample → Sample Z to get sample into approximate focus (in 0.5 to 1μm steps) if needed

  - Motion control, go to sample theta= –30

  - Use line tool to mark a distinct and tractable fine feature

  - Motion control, go to sample theta= +30

  - Use line tool to mark the same point

  - Motion Control → Sample → Sample Z; use steps of 0.2 to 1μm to move the same feature perfectly between the two lines

  - Motion control → Sample theta = 0

  - Acquisition settings: Abort and delete image

- Focus the objective (zoneplate)

  - Start acquisition (gears button) → Mode continuous, 0.5s exposure

  - Move Motion Controller → Zoneplate → Zoneplate Z to get sample into precise focus (0.5 to 2μm steps ONLY)

  - Focus image on cell features of interest, not support film or fiducials

- Alternative focussing method

  - Acquisition settings: mode focal series = camera setting 0.5s, binning 1; Focal series = Images per projection 1, wait between images 0.2, Focal axis Zoneplate Z, Range 10, Start

  - Save as FS_<IncreasingNumber>

  - Choose best focal slice and remain on that slice

  - Go to the forward arrow button ‘>’ on the right of the tilt series and select Positions → Note down the Zoneplate Z Actual position

  - Go to the motion controller → Zoneplate → Change Zoneplate Z to the recorded value and click GO

  - Check that the sample has moved to the desired Z coordinates by running a continuous acquisition.

- Test maximum rotation (see above critical warning)

  - Acquisition settings: mode continuous, camera setting 0.5s, start

  - Motion Control → Sample → Sample theta step =5, Sample theta =50

  - Step sample theta down to –55 and up to +55. Record extremes with clear visibility (grid bars or sample density can obstruct light at high tilts).

  - Note pixel counts at high tilts and consider increasing exposure if counts are low (<3,000)

  - Acquisition settings: Abort and delete image

- Decide on data collection strategy (depends on sample stability and user project needs)

Conservative: 1deg steps, 0.2 delay (i.e. -55 to +55: 111 images)

Standard: 0.5deg steps, 0.2 delay (i.e. -55 to +55: 221 images)

Exhaustive: 0.2 deg steps, 0.2 delay (i.e. –55 to +55: 551 images)

- Collect tilt series

  - Acquisition settings:

  - Acquisition Modes: Tomography

  - Camera Settings: 0.5-5s, binning 1

  - Sample Settings: tick pixel size = 0.01

  - Tomography: images per projection 1, wait between images 0.2, total projections <YourTotalNumberOfImages>, Angle Start= –55, Angle End= +55, Return to zero degrees when finished = ticked, collect reference none, variable exposure none, start

  - Save as <TiltSeriesNumber>_Tomo_<GridBox>_<GridStoragePosition>_<GridArea>_<NumberOfFOV> (ie 01_Tomo_K1_G1_D4A_ROI1)

  - When finished, click on arrow top-right at Image side panel → Reference: load single reference, Shifts: uncheck Dynamic runout

  - Save Tilt Series

2D. XMBatch acquisition

XMBatch works by importing single image files (snapshots) that contain motor positions and exposure settings to allow scans to be quickly queued and run in a batch.

To use XMBatch, follow these steps:

- Create a folder for the batch.

- Add scan snapshots and a reference snapshot (see the file naming convention below). Refence snapshots can be reused if the reference location has not changed – to do this move or copy the reference file into the new batch folder.

- Select ‘Load Scans from Folder’ to import files. This will import all found scan snapshot files and the most recent reference snapshot file in a folder. Please note:

  - Files will not be visible when selecting the folder.

  - XMBatch does not remember what has been used, so it will reimport previously used snapshots if they are in the selected folder – prepare your batch folder carefully!

- Check the exposure and angle settings.

- If you are happy with the settings, press ‘Execute Scans’.

Snapshot file naming

- Scan snapshots:

  - Must include ‘snapshot’ in the file name (not case sensitive).

  - Tomo angles can be specified in the name in the format ‘<min angle>to<max angle>@<step size>’, e.g. ‘-55to+55@0.5’. If you don’t include the -/+ signs, XMBatch will assume <min angle> is negative and <max angle> is positive.

  - If no angle information is given, the XMBatch default values will be used. This can be edited in the GUI before executing the scan.

- Reference snapshots:

  - Must include ‘refsnap’ in the file name (not case sensitive).

  - If multiple ‘refsnap’ files are found, the most recently modified file be used.

  - Refsnap files can be reused by ensuring it is in the imported folder. Check that the position is still valid.

Any files not named with the above convention can be imported separately using the “Append Single Scan” button for scans or “Load Reference File” button for the reference.

Quick end of beamtime shift checklist

- Change front end (FE) status to partially close

- Warm up Pixis camera to 23°C