B07-C TCUP AP-XPS User Guide - TCUP

January 2020

Table of Contents

1.	User Induction Checklist	2

2.	TCUP Venting Procedure	3

3.	TCUP Pumping Procedure	4

4.	Changing Photon Energy	5

5.	Sample position	6

6.	Initial Alignment via SPECS Prodigy	7

7.	XPS Measurements	8

8.	NEXAFS Measurements	9

9.	Normalization	10

10.	Au Foil Standard	10

11.	Beamline Transmission Curves	11

12.	Introducing Gases (Manual Control)	12

13.	Pumping Out Chamber	13

14.	Eurotherm Operation	14

15.	Troubleshooting – GDA acquisition problems	15

16.	Troubleshooting – PGM Motion Failure	17

17.	Troubleshooting – Magnetic Field	18

18.	Useful GDA Commands	19

20.	Beamsize at the sample on B07C	22

21.	Au4f spectra under H2 and N2	23

22.	Data analysis	25

23.	Protocol for repairing disconnected TCUP Xp motor (broken cable tie)	28

24.	Installing the HDF5 package for Igor	29

# User Induction Checklist

BEAMLINE INTRO

- Walkthrough

- Beamline staff – introduce, explain roles, how to contact

- LOGIN @ users account in diamond ws-001 computer

- Beamline user guide –TCup manual

- Food/drink rules – control room only, only hand-held stuff

- Restaurant/kitchens – opening hours

  - Canteen

Weekdays: Breakfast 7.30am – 9.30am.

Evening meal 5.30pm – 8.00pm (last orders at 7.45pm).

Weekends: Breakfast 8.00am – 9.00am.

A packed lunch can be pre-ordered. Orders must be placed by 10 am.

  - Chilton town (~2-3 miles) The Rose & Crown/Main St, Didcot OX11 0RZ

- Remove waste/samples at end of beamtime

SAFETY

- Fire escapes – locations and muster point

- Fire extinguishers – locations, only use if safe/trained

- Gas alarms - procedure

- EHCs – 8787

- Control Room – 8899

- Specific hazards – magnetic fields, high voltages,

- Lone working procedures – notify EHCs

- Motion stops

- Beamline power stops

EXPERIMENTAL

- Vacuum schematic

- EPICS – interlocks, valves, shutter, PGM, mirrors

- GDA – perspectives, running/saving scripts

- Data retrieval/analysis  – FTP, DAWN, Igor

- Cryogens – PPE

- Gas bottles – EHCs to connect/replace

- Sample prep – no chemicals, nanomaterials

- Lab 51 – induction & lab form required

- Sign form on door

# TCUP Venting Procedure

- Close Experimental Shutter SHTRC

- Close Ea shutter ESHTR (if needed)

- Move sample manipulator back to minimum Xp position (NAP-ES  SM2(TCUP))

- Insert blocking bit into Xp motor

- Disconnect the software and turn off QC (blue box) and power supplies.

- Open GV13

- Close V12

- Close V22

- Close V25

- Turn off RGA 21 and RGA 25 (Hiden) (if needed)

- Turn off HCG-25 emission (in the rack)

- Check V49 is open

- Close V35

- Turbo 24 and 25 will automatically slow down and stop. Will take few minutes to completely stop.

- Close V24

- Close V48

- Open N2 bottle and set regulator to gentle flow

- Switch key to service mode on HMI

- Open V47 (wait for 5 min or until turbo 24 and 25 are stopped, red color in HMI)

- Monitor PIRG25 to check when chamber is vented

# TCUP Pumping Procedure

- Close V47

- Close N2 bottle and regulator

- Check blocking bit is in Xp motor

- Check V49 is open

- Close V46

- Open V35

- Open V48

- Wait until PIRG 24, GPIRG 33, CAP-05 are all green

- Start turbo 24 and turbo 25

- Close V48

- Open V46

- Start turbo 30

- Open V24 (this will no open until turbo 30 is green)

- Turn on HCG-25 emission

- Switch key to ‘Normal’ mode on HMI

- Remove Xp motor blocking bit

BEAM ON SAMPLE AFTER VENTING

Open V22 and V25

Open B07CNAP-ES vacuum synoptic  interlock  Reset

Reset V12

Open V12

Reset Exp Shutter SHTRC

Open Exp Shutter SHTRC

Open EShutter (if closed)

Switch on analyser power supplies and QC (blue box)

Then, you align the sample.

# Changing Photon Energy

PGM  Energy  override

- No override: slow movements – for NEXAFS

- Override : fast movements - XPS

Set the desired energy on the PGMc EPICS screen

Check cff = 2

View D5c_cam2 diagnostic (camera on top of exit slits) to see image of beam centred on the cross (cross at x:540, y:840). It should appear as a “flame” as shown below:

# Sample position

Camera view

Mark on flange

Actual view

# Initial Alignment via SPECS Prodigy

Connect the manipulator wires

Select desired photon energy as described in Section 4

Check sample manipulator NAP-ES  SM2 TCUP

Check that valves 22, 25, 12 and ESHTR (entry shutter) are OPEN.

Reset NAPES vacuum interlocks before opening valve 12

Check that ea_shutter (ESHTR fast shutter in beamline entrance) is OPEN

Open SHTRC [experimental shutter]

Switch on analyser power supplies and QC (blue box)

Go to SPECS Prodigy via remote desktop (Diamond Launcher  SPECS Analyser PC)

Click Phoibos (magnetic field compensation) @ Device Control

Start fixed energy scan for alignment (check KE, photon energy) @ experiment editor

Move Xp stepwise, (0.01 mm steps) monitoring D22c camera and maximise counts in analyser

Tweak  M4c_pitch and M5c_pitch to maximise counts

Stop the scan running in SpecsProdigy

Important – if you cannot see the laser dot then the sample may be in contact with the analyser cone – STOP approaching immediately.

CAMERAS:

NAP-ES  camera  D22C  set-up  exposure…

# XPS Measurements

Pre-scan Checks:

- SPECS config file: EPICS_magnetic_field_20171020 (This sets the autocompact to compress both energy and angular channels) @ remote control in SpecsProdigy (unclick “allow” box  change  click “allow” box again)

- Ensure “Allow remote connections” is ticked in SPECS Prodigy

- Sample is grounded either through FEMTO or via BNC shorting cap. Depending on the sample holder and the experiment the grounding can be either through the drain current connections or through the thermocouple.

- In some cases, charging has been observed when thermocouple is connected.

- PGM energy: OVERRIDE

Example XPS script:

(set up and save name_of_sequence.seq with desired regions etc before sending)

Sequence name

Beamtime

directory

analyser.setSequenceFile('/dls/b07-1/data/2018/si1234-1/xml/name_of_sequence.seq')

scan dummy a 0 0 1 analyser ring_current

PGM energy override

# NEXAFS Measurements

Pre-scan Checks:

- PGM energy: NO OVERRIDE

- SPECS config file: EPICS_magnetic_field_20171018 (This sets the autocompact to off for both energy and angular channels) @ remote control in SpecsProdigy (unclick “allow” box  change  click “allow” box again)

- Thermocouple is disconnected

- Heater is disconnected

- Sample drain current is connected to FEMTO 23

- D21c is in the “Mesh” position (GDA command: pos d21c_positioner ‘New Mesh’)

- Ensure“Allow remote connections” is ticked in SPECS Prodigy

Example TEY script:

End energy

Acquisition time (sec)

scan pgm_energy 500 530 0.5 sample_current 2 femto_21 2 ring_current

Acquisition time (sec)

Energy

step

Start

energy

Example AEY script:

(set up and save name_of_sequence.seq with desired settings for snapshot mode before sending)

Snapshot name

Beamtime

directory

analyser.setSequenceFile('/dls/b07-1/data/2018/si1234-1/xml/name_of_sequence.seq')

scan pgm_energy 280 305 0.05 analyser sample_current 5 femto_21 5 ring_current

PGM energy NO override

Femto_21  Io measurement by mesh (diagnostic 21)

NAP-ES  motion  ‘mesh’ or ‘out of beam’

# Normalization

TOP-UP normalization:

- ring_current

- S2c as detector GDA Command (s2c_ytop_current)

# Au Foil Standard

Counts during approach

hv:

Ekin:

Epass:

T: 1s

Slits:

Au 4f XPS

# Beamline Transmission Curves

# Introducing Gases (Manual Control)

Open V36 and V50

Open leak valve until pressure starts to rise on HCG-25

Turn OFF emission on HCG-25

Close V24

Close V46

Close V49

Open V48

Operate V37 – manual butterfly valve to regulate pressure

See desired pressure on CAP-03

# Pumping Out Chamber

Close the leak valve(s)

Close Manual valve 37

Open valve 49

Close V48

Open V46

Restart Turbo 30

Open V24

Wait until Turbo 24 is on

Turn on HCG 25 back

# Eurotherm Operation

NAP-ES  Motion  Sample Heater

RETRACT SAMPLE BEFORE HEATING!

The eurotherm can be used either manually, through its EPICS control window, or via GDA.

Pre-heating Checks:

- Ensure that the heater cable is connected at the manipulator

- Ensure that the thermocouple (green cable) is connected at the manipulator and a sensible temperature is being read by the Eurotherm

- Ensure the power supply (above Eurotherm) is switched ON

- Ensure the sample has been retracted away from the analyser cone

- Compressed air to be flowed through the manipulator

Heater EPICS Operation (located at NAPES  Motion  Sample heater):

- Type desired heating rate and temperature set point

- When at stable temperature, re-approach analyser cone with fixed energy scan

- For more rapid cooling after heating experiment, ask for compressed air to be flowed through the manipulator – see beamline staff

GDA script that records XPS at 10° increments from 100-400 °C:

analyser.setSequenceFile('/dls/b07-1/data/2017/siXXXX-1/xml/name_of_sequence.seq')

scan sample_temperature 100 400 10 analyser

Other Commands:

To know the current ramp → sample_temperature.getRampRate()

To select/change a 5 °C/min ramp → sample_temperature.setRampRate(5)

To set the new set point → pos sample_temperature XXX

To stop the ramp before raise the final temperature → sample_temperature.stop()

To set the accuracy level → sample_temperature.setAccuracy(1.0)

To get the current accuracy level → sample_temperature.getAccuracy()

# Troubleshooting – GDA acquisition problems

If the scan fails to run – first check that the analyser EPICS interface is still connected: Go to the synoptic and select NAP-ES:SPECS Analyser and check that the CCD is connected (if not, just click reconnect)

Connection status

The connection between GDA and the analyser occasionally drops out, which results in slow/faulty updating of the spectra within the GDA window. To recover, in the first instance try the following:

Dump current plot data (save to nexus file)

Close the GDA client

Restart GDA server (Diamond Launcher  Start B07-1 GDA Servers)

Wait until this process finishes (generally a couple of minutes)

Restart GDA client (Diamond Launcher  Start B07-1 GDA Client)

In case this doesn’t resolve the problem, a reset of the IOC may be required, as detailed below:

Close the GDA client

Go to “Hardware Status” on synoptic screen (lower left)

Click Start/Stop tab for NAPES (Analyser) BL07C_EA_IOC_02

Restart this IOC

Restart GDA server (Diamond Launcher  Start B07-1 GDA Servers)

Restart GDA client (Diamond Launcher  Start B07-1 GDA Client)

Reconnect analyser in EPICS window

# Troubleshooting – PGM Motion Failure

An amplifier loss on the Geobrick that controls the PGM can be recovered as follows:

- By pressing STOP on the energy window of the PGM screen and inputting the desired value again.

- Hardware status  Geobrick PGMc  Terminal  Connect  type: ena plc6  Send  Disconnect

# Troubleshooting – Magnetic Field

- Disconnect the device in Specs prodigy

- Switch off the three power supplies of the coils

- Unplug and re-plug the Armin plug into the socket. It is in the rear part of the analysis rack   3, bottom right, and labelled.

- Switch on the three power supplies

- Connect the device in Specs prodigy

- Check if the enable LED is on in the three power supplies

# Useful GDA Commands

| GDA Command (italics = stuff you can change) | What it does |
| --- | --- |
| analyser.getWorkFunction() | Gets current value of workfunction |
| analyser.setWorkFunction(5) | Sets value of workfunction |
| femto_name.setSettleTime(500) | Sets settle time of femto_name |
| femto_name.setAutoGain(True/False) | Sets autogain on/off for femto_name |
| femto_name.isAutoGain() | Gives autogain setting for femto_name |
| pos expt_shutter 'Close' | Closes experimental shutter (replace Close-Open) |
| pos pgm_energy 743.3 | Sets photon energy to 743.3 eV |

Femto Names

| sample_current | Sample drain current (Femto 23) |
| --- | --- |
| ap1c_current | BLE aperture (Femto 22) |
| d21c_current | Diagnostic 21 (Femto 21) |
| d7c_cell | Gas cell |
| d7c_diode | Gas cell photodiode |

Acquiring XPS spectrum via script

analyser.setSequenceFile('/dls/b07-1/data/2018/siXXXX-1/xml/name_of_sequence.seq')

scan dummy a 0 0 1 analyser

Acquiring NEXAFS spectrum via script

pos ea_shutter ‘Close’

pos pgm_energy 280

pos d21c_positioner ‘Mesh’

pos ea_shutter ‘Open

analyser.setSequenceFile('/dls/b07-1/data/2018/siXXXX-1/xml/name_of_sequence.seq')

scan pgm_energy 280 305 0.05 analyser sample_current 5 femto_21 5

pos ea_shutter ‘Close’

pos d21c_positioner ‘Out of Beam’

The above scans the photon energy from 280 eV to 305 eV, in 0.05 eV steps, acquiring AES NEXAFS via the analyser (assuming snapshot mode), TEY from the drain current, and an I0 from the mesh on d21c.

Changing Grating via GDA --> all need to be checked

#400l/mm grating, mirror: Pt stripe

pos pgm_grating '400 lines/mm'

pos pgm_mirror 'Pt Stripe'

#400l/mm grating, mirror: Rh stripe

pos pgm_grating '400 lines/mm'

pos pgm_mirror 'Rh Stripe'

# 600l/mm grating: Rh stripe, mirror: Rh stripe

pos pgm_grating '600 lines/mm'

pos pgm_grating_x -91.7

pos pgm_grating_x -91.9

pos pgm_mirror 'Rh Stripe'

# 600l/mm grating: Rh stripe, mirror: Pt stripe

pos pgm_grating '600 lines/mm'

pos pgm_grating_x -91.7

pos pgm_grating_x -91.9

pos pgm_mirror 'Pt Stripe'

# 600l/mm grating: Pt stripe, mirror: Pt stripe

pos pgm_grating '600 lines/mm'

pos pgm_grating_x -107.1

pos pgm_grating_x -106.9

pos pgm_mirror 'Pt Stripe'

# 600l/mm grating: Pt stripe, mirror: Rh stripe

pos pgm_grating '600 lines/mm'

pos pgm_grating_x -107.1

pos pgm_grating_x -106.9

pos pgm_mirror 'Rh Stripe'

- Data Visualization/Extraction

GDA: 		Script perspective

Specs perspective

DataVis

DAWN:  (module load dawn\2.5   dawn  @ October 2017)

Trace or DataVis perspective to visualize and compare:

File  new project  Data  Data project (linked…)  select the external folder

Data extraction:

Windows PC  explorer: \\data.diamond.ac.uk\b07-1\data\2019

FTP:	Load your data from your explorer at: \\diamftpserv01\ftpanon\B07

Take the data from:  ftp://ftpanon.diamond.ac.uk/B07/

Dawn:  Convert .nxs to XY

Processing Perspective

Select a file		dataset: spectrum

line [1d]

select KE, BE..

Processing window		export to text file

Run from data slice view window

To convert to .txt files:

Within the processing folder of your own experiment folder (typically dls/b07-1/data/2019/sixxxxx/) create a new folder named convert in processing and copy the data files to that folder.

Then from a terminal input the command:

/dls_sw/b07-1/scripts/b07_convert_nxs_to_txt_new.py /dls/b07-1/data/2019/sixxxx/processing/convert

- Data extracted: kinetic_energy vs intensity

/dls_sw/b07-1/scripts/b07_convert_nxs_to_txt_new2.py /dls/b07-1/data/2019/sixxxx/processing/convert

- Data extracted: binding_energy vs intensity

This will makes .nxs files in the ‘convert’ folder to .txt files

# Beamsize at the sample on B07C

Measurement Procedure:

PGM: 600 l/mm; Pt mirror
Photon energy: 1200 eV

Sample: thin rectangular mica sheet (<<0.1 mm x 11 mm x 7 mm) on bare sample holder; edges aligned with yp/zp directions of sample manipulator. The sample holder is at normal incidence.

| y gap size / mm | vertical beam size 12% - 88% / mm |
| --- | --- |
| 0.200 | 0.103 |
| 0.150 | 0.090 |
| 0.100 | 0.075 |
| 0.075 | 0.070 |
| 0.050 | 0.066 |
| 0.025 | 0.062 / 0.065 |

| z gap size / mm | vertical beam size 12% - 88% / mm |
| --- | --- |
| 0.173 | 0.090 |
| 0.185 | 0.090 |

# Au4f spectra under H2 and N2

h = 500 eV  S4c = 0.8 x 0.05

h = 500 eV  S4c = 0.8 x 0.05

h = 500 eV  S4c = 0.8 x 0.025

# Data analysis

Users can analyse their data in the beamline hutch windows computer with Igor pro software.

Copy/access your data.

- Make a local folder to copy your data.

- If you are logged in with your credential (diamond login id and password),

Then type \\data.diamond.ac.uk  in the address bar to access your data from

B07-1>data>2019>your data directory (for example si2218184-1)

- If you are logged in as b071user then you have to access your data from the same directory through ftp via Filezilla ftp client software

Your data folder path: dls/b07-1/data/2019/SIxxxx

Put your diamond username

Igor procedure files

We have two procedure macros for igor, which enables us to easily convert nexus files into igor waves and analyse xps data.

- In Igor pro macros HDF Files allow you to convert nexus data files to igor binary waves

- XPS macro allows you to perform some analysis like normalization, background subtraction, fermi energy correction etc.

- If you are logged in as b071user

The macros are already loaded.

- If you are logged in as yourself (your user name)

First check if you have hdf browser installed in igor.

If yes, then load the macros directly from.\\diamftpserv01\ftpanon

B07-->Igor macros

If you do not see hdf browser installed. First install the hdf browser and then load macros.

# Protocol for repairing disconnected TCUP Xp motor (broken cable tie)

Situation: The plastic cable tie that connects the Xp motor drive to the manipulator has broken and therefore you cannot move the manipulator – this is generally caused by leaving the blocking bit in place and moving the Xp motor by mistake.

Protocol:

- Vent the chamber

- Take out the manipulator

- Blank off the manipulator port (KF50)

- Manually rotate the metallic cylinder piece at the end of the motor lead screw to a horizontal position (parallel to the metallic bit where the cable tie is located) - see picture below

- Replace the small plastic tie for a new one and leave loose

- Close v49 and the manual control valve (V37)

- Open v48, to pump by the scroll (bypassing turbo 29)

- Open manual valve (V37) slowly and check/observe how the bellows contract - see picture

- Once the cylinder and the rectangular piece are together, tighten the cable tie and snip end off

- Continue pumping (open to turbo 29 when appropriate)

- Test all range of motion and home Xp if needed

- Good luck!

# Installing the HDF5 package for Igor

Installing HDF5 requires activating two files – the HDF5 XOP file and the “HDF5 Browser” procedure file – by placing aliases (Mac OS X) or shortcuts (Windows) for them in special folders.

A. From the Help menu, choose “show Igor Pro User Files”. This displays the Igor Pro User Files folder on the desktop. These instructions assume that you are using the 64-bit version of Igor.

From the Help menu, choose “Show Igor Pro Folder”. This displays the Igor Pro Folder on the desktop.

Make an alias (Mac OS x) or shortcut (Windows) for the following file: Igor Pro Folder\More Extensions (64-bit)\File Loaders\HDF5-64.xop

Move the alias or shortcut into “Igor Pro User Files\Igor Extensions (64-bit)”. This activates the HDF5 XOP, which adds operations to Igor for interaction with HDF5 files.

The HDF5-64.xop file contains within it the necessary HDF5 libraries so you do not need to install libraries from the HDF Group.

B. Make an alias (Mac OS X) or shortcut (Windows) for the following file:

Igor Pro Folder\WaceMetrics Procedure\File Input Output\HDF5 Browser.ipf

Move the alias or shortcut into the “Igor Pro User Files\Igor Procedures” folder. This activates the procedure file that implements an HDF5 browser and adds a “New HDF5 Browser” menu item in the Data  Load Waves menu.

C. Restart Igor Pro and then read the HDF5 Guided Tour section below.

To uninstall the HDF5 package, delete the aliases or shortcuts created above and restart Igor.

For background information on special folders such as the Igor Extensions and Igor Procedures folders, see Special Folders.